Gen5 under analog fully conditioned mode in VICON NEXUS

- Gen 5 Setup:

Digital Outputs on ‘English Units’

Analog Qutputs on ‘Full Condition’

Then click on the configure button corresponding to the amplifier to be setup.

- Amplifier Configuration:

‘Excitation Voltages’ on 5V. Other voltages could be selected. 
Under , you will find the section ‘Analog Sensitivities’ in which you have a number that corresponds to the analog output sensitivity. Changing that sensitivity will update the analog output range in the table above and optimize the output resolution. Make sure the analog output range corresponds to your measuring range and close to the analog input range.

- Button to Open the AMTI calibration matrix.

- Calibration Matrix:
Enter the main diagonal terms in that matrix according to the number entered in the ‘Analog sensitivities’ section of the amplifier setup.

- Change the correction to ‘1000’ . Because the analog sensitivities of the Gene 5 are in mV, you have to convert them in V for Nexus.

- Warning
In items 1-5, the platform origin offsets are not computed. This needs to be entered separately in Nexus.
You will find that section on the left under the Force Plates → Properties → Advanced→ Show Origin

Pictures and setup description provided by PROPHYSICS CH.

# Text boxes

==========3

4

1

2

5